## Slide 1: Exercises – biological signals
Exercise 5 - SS 2014 – Michel Kana

## Slide 2: What will we do today?
The physiology of Respiration
Structure of the Respiration Signal
Respiration measurement with BIOPAC
Summary

## Slide 3: The physiology of Respiration
The respiration system consists of structures involved in ventilation and gas exchange
  Airways that lead air from the external environment to the exchange surface of lungs
  Alveoli that form an exchange surface where oxygen moves to blood and carbon dioxide moves from blood.
  Bones and muscles of the thorax and the abdomen that assist in ventilation.
The primary functions of the respiration system include
  Gas exchange
  Homeostatic regulation of body pH
  Protection from inhaled pathogens
  Vocalization

## Slide 4: The physiology of Respiration
Breathing is an active process that uses muscle contraction to create a pressure gradient.
  When the diaphragm contracts, it drops down toward the abdomen about 1.5 cm, increasing thoracic volume.
  When costal and scalene muscles contract, they pull the ribs upward and out, increasing thoracic volume.
The basic rhythm of breathing is established by inspiratory and expiratory respiratory centers in the medulla.
The air moved during breathing is divided into four lung volumes.
  Tidal volume is the volume of air that moves during a single inspiration or expiration (ca. 500 mL).
  Inspiratory reserved volume is the additional volume of air that can be inhaled after tidal volume is reached during inspiration (ca. 3000 mL).
  Expiratory reserved volume is the additional volume air that can be exhaled after the tidal volume is reached during expiration (ca. 1100 mL).
  Residual volume if the volume of air that remains in the lungs after a maximal exhalation (ca. 1200 mL).

## Slide 5: Structure of the Respiration Signal
Respiration creates change in thoracic or abdominal circumference.
The measurement of the thoracic or abdominal circumference as a voltage is called respiration signal.
The respiration signal contains various information:
  The respiration rate: number of inhalation/exhalation cycles per minute (ca. 12-14).
  The tidal volume: estimated from the amplitude of the respiration signal during quiet breathing. Calibration is required.
  The inspiratory reserved volume: estimated from the amplitude of the respiration signal during forced inhalation.
  The expiratory reserved volume: estimated from the amplitude of the respiration signal during forced exhalation.
  The inspiration and Expiration duration

## Slide 6: Structure of the Air Flow Signal
Respiration creates bi-directional airflow.
The measurement of airflow as a voltage is called airflow signal.
Volume measurements are obtained by integrating the airflow signal.

## Slide 7: Exercice 1: Respiration measurement with BIOPAC
Biopac MP35 measurement system
  Respiration is recorded using the Biopac SS5LB transducer plugged in the first channel.
  Air flow is measured with a Biopac SS11LA transducer plugged in the second channel.
  Air temperature is measured with a Biopac SS6L transducer plugged in the third channel.
Biopac Student Lab PRO software
  The acquisition is set up at a sampling rate of 200 Hz.
  Analog Channel CH1 have the preset Respiration (SS5BL)
  Analog Channel CH2 should have should the preset Airflow (SS11L)
  Analog Channel CH3 should have should the preset Temperature (deg C)
  Calculation Channel C1 should have the preset Respiration Rate (from CH1)
  Calculation Channel C2 should have the preset Lung Volume (from CH2)
Respiration parameters calculation
  Estimate the breathing rate, the mean tidal volume, the mean inspiration and expiration duration

## Slide 8: Exercice 2: Respiratory volume estimation
Procedure
  Subject is instrumented for Respiration measurement with Biopac.
  Subject should perform a normal inhalation, followed by a forced inhalation to maximal lung volume during 60 sec.
  Subject should perform a normal exhalation, followed by a forced exhalation to minimal lung volume during 60 sec.
Evaluation
  Calculate the tidal volume during forced exhalation and forced inhalation.
  Compare the obtained values with tidal volume during quiet breathing and estimate the expiratory and inspiratory reserved volumes.

## Slide 9: Exercice 3: Respiratory patterns
Procedure
  Subject is instrumented for Respiration measurement with Biopac.
  Subject should perform 6 cycles deep breathing during 60 sec.
  Subject should then perform 60 cycles hyperventilation during 60 sec.
  Subject should then record a couple of respiration cycles during cough and yawning.
Evaluation
  Calculate the mean tidal volume, mean inspiration and expiration durations for each experiment.
  Compare the values between experiments.

## Slide 10: Exercice 4: Breathe holding
Procedure
  Subject is instrumented for Respiration and pulse measurement with Biopac.
  Subject should record 60 sec of quiet breathing followed by 10-30 sec breathe holding.
  Subject should resume 60 sec of quiet breathing.
Evaluation
  Calculate the mean heart rate, respiration rate, mean tidal volume, mean inspiration and expiration durations for each of the three phase.
  Compare and discuss the values.